HiTrap Q Ion Exchange Column

- Adjust the buffer of the sample to match the Start Buffer.  Dialyze in Tris Dialysis Buffer overnight or exchange buffers using a concentrator.

- Turn on UV lamp.  Turn on computer.  Go to FPLC Director.  Log on as “allison”.  Make Start Buffer and Elution Buffer.  Set up tubes.  Move the arm so that its mark is just in front of the first tube.  Make gels.

- Obtain HiTrap Q column.  Notes for dealing with column:  Never let the column get air in it.  Therefore, for each wash, fill the neck of the column with buffer, attach the adaptor, fill the neck of the adaptor with buffer, fill the syringe, flick out all air bubbles, attach the syringe, and inject the solution at a pace where it slowly drips. Keep the column upright.

- Wash preservatives out of column.  Using syringe, apply 25ml Start Buffer.

- Remove any left-over proteins.  Apply 25ml Elution Buffer.

- Equilibrate column.  Apply 25ml Start Buffer.

- In the FPLC Director, go to “Run/Control”.

- Wash system with water.

  - Put Buffer A and Buffer B feeder tubes in dH2O.

  - Manual(Pump(Wash(Waste(Check A and B(Start Wash

- Wash system with buffers.

  - Put Buffer A tube in Start Buffer.

  - Put Buffer B tube in Elution Buffer.

  - Manual(Pump(Wash(Waste(Check A and B(Start Wash

- Rinse loop.

  - Manual(Flow Path(Valves(New Position-Load(Set(Exit

  - Fill syringe with needle with Start Buffer.  Knock out bubbles.  Inject into little hole.  Should fill loop.  Repeat.  Leave syringe in.

- Install column in cold cabinet.

  - Unscrew adaptor.

  - Attach bottom of column to white box.

  - Fill top of column with Start Buffer and attach top feeder tube.

  - Put paper towels underneath to see if the column is dripping.

- Wash column with Start Buffer.

  - Manual(Pump(Flow A+B(2ml(OK

  - Run until blue UV280 line goes down (~10min).

- Auto zero.  Press button until red light comes on.

- Draw ~7ml of dialyzed protein into syringe with needle.  Remove bubbles and inject into little hole.  Leave syringe in.

- Run “hitrapq” method.  File(Run Method(hitrapq(Run.

  - The “hitrapq” method starts the fraction collector and injects the protein onto the column.  It then applies about 45mls of Start Buffer to allow your protein to bind to the column and to wash out the contaminating proteins that do not bind.  It then slowly increases the amount of salt in the buffer by increasing the percentage of Elution Buffer from 0% to 100%.

- Go to Main Menu (at bottom of screen).  Evaluate which tubes have protein by looking at chrom1 for the file.

- Run appropriate samples and 1 random sample.

  - 15(l sample + 3(l 5X RSB.  Boil at 100(C for 5min.  Run 10(l.

  - Stain with Coomassie Blue for ~25min.  Destain and rinse with dH2O.

  - Dry gel to save it.

- Pool tubes that look good (those that are at the peak) into a 50ml conical on ice.  Discard other tubes.

- Dialyze.  (Gets rid of salt in sample).

  - Cut ~9in of dialysis tubing (in cold cabinet—keep moist).

  - Soak in dH2O for a few seconds.

  - Clamp on one end.

  - Fill with sample.

  - Clamp top ~1in from the top of the sample.

  - Put in PBS Dialysis Buffer at 4(C overnight with slow spinning.

- Detach the column.

- Remove any left-over proteins.  Apply 25ml Elution Buffer.

- Regenerate the column.  Apply 25ml Start Buffer.  Column is now ready to use again.

- Store column.

  - Apply 5ml dH2O.

  - Apply 25ml 20% ethanol.

  - Seal the column.  Store at 4ºC.

- Turn lamp off.

- Wash system with water (can be done next day).

  - Put Buffer A and Buffer B feeder tubes in dH2O.

  - Manual(Pump(Wash(Waste(Check A and B(Start Wash

- Wash system with ethanol (can be done next day).

  - Put Buffer A and Buffer B feeder tubes in 20% Ethanol.

  - Manual(Pump(Wash(Waste(Check A and B(Start Wash

- Concentrate protein.

  - Pour dialyzed protein into a 50ml conical on ice.

  - Get concentrator (in microscope room—Millipore Amicon Ultra).

  - Fill inside of container with protein sample.  Screw on top.

  - Spin at 3,000rpm for 10min at 4(C.  (Small things will cross filter and go into the exterior of the chamber).

  - Discard the liquid in the exterior chamber.

  - Repeat as needed, adding liquid to the interior chamber and spinning until sample is gone.  Do not concentrate too much or protein will precipitate.

- Bradford Assay

  - Samples to read

    - Blank (800(l dH2O + 200(l BioRad)

    - BSA Standards (800(l dH2O + 2, 4, 6, or 8(l 1:10 BSA + 200(l BioRad)

    - Importin ( samples (800(l dH2O + 2 or 4(l sample + 200(l BioRad)

  - Put 800(l dH2O in cuvettes.

  - Dilute BSA 1:10 to a final concentration of 1mg/1ml.

  - Pipette BSA into cuvettes.  Do not change tip.

  - Pipette samples into cuvettes.  Do not change tip.

  - Vortex cuvettes.

  - Add 200(l of BioRad Protein Assay Reagent.  Vortex.

  - Measure OD at 595.

  - Use Cricket Graph to find concentration.

    - Enter data for standards

    - Graph(New Graph(Line    y=mg/ml    x=OD

    - Options(Curve Fit(method(linear)    .98<r<1

    - Use equation to calculate concentration of sample.  Plug in the average of the readings for the sample for x.  y=conc.

- If concentration is too low, concentrate again.

  - May have to use small concentrator.

  - Apply liquid to the top.

  - Spin at 3,000rpm for 30min at 4°C.

  - Keep liquid on top.  Use a transfer pipette to squirt the membrane and free protein.  Save.

  - Repeat as necessary.  Find concentration.

- Add 10% glycerol, aliquot, store at -80(C.

Materials

1L 20mM Tris pH 7.8

2.42g Tris Base Ultrapure

1L dH2O

-pH to 7.8

Tris Dialysis Buffer

4L 20mM Tris pH 7.8

1.23g DTT

Start Buffer = 20mM Tris pH 7.8 + 2mM DTT

1L 20mM Tris pH 7.8

0.31g DTT

-filter sterilize

Elution Buffer = 20mM Tris pH 7.8 + 2mM DTT + 1M NaCl

1L 20mM Tris pH 7.8

0.31g DTT

58.44g NaCl

-filter sterilize

PBS Dialysis Buffer

400ml 10x PBS

1.23g DTT

3600ml dH2O

10x PBS

80g NaCl

2g KCl

21.8g Na2HPO4·7H2O

-add dH2O to a final volume of 1L

-pH 7.2-7.4

-Filter sterilize